## Slide 1
Information of Device
- Product name: Classic (POS, UTL, ANT), Perfit ZR (HT, ST, UT)
- Model name: Total 4830 types including model CP 4019 white C
- ISO6872 Type: 2, Class: 5 (Classic POS, Perfit ZR HT), 4(b) (Classic UTL&ANT, Perfit ZR ST&UT)
- Color: white and shade according to VITA shade guide (classic) / Size: 4019, 5519, B65, 10-30mm
- Target population: All ages excluding newborns.
- It can be used in the following equipment. (representative)
; Imes-icore GmbH: CORiTEC series (150i, 250i, 350i) / Amann Girrbach AG: Ceramill series
; Roland: DWX-50, DWX-51D, DWX-52D, DWX-52DCi / vhf CAMFACTURE AG: K4, K5, K5+, S5
- Conventional cementation with zinc oxide phosphate cement or glazor cement is recommended
to mount the prosthesis. Luting composites are also available.
- It is recommended using diamond coated burs for precise use.
- The types of cement recommended for zirconia restoration is RMGI and Self adhesive resin cement
Indication
Anterior and posterior dental restorations such as crowns and bridges. Or prosthetic treatment due
to tooth loss and caries
Intended Purpose
Designed by professional dental technicians to create custom all-ceramic restorations. Use in CAD /
CAM milling systems to make dental restorations such as crowns and bridges. It can be used to
create anterior and posterior dental restorations.
Adverse reaction
hypersensitivity to cold, gingival bleeding, pain, chipping of the veneering porcelain, loss of
cementation, localized gingival irritation, loss of retention, infection, allergic reactions, irritation,
Side effects
postoperative tooth sensitivity, framework fracture, temporary pain
General handling instruction (Preparation before use)
- Visually check the product for damage, presence of impurities, and abnormalities.
- Make sure that the appearance and shape of the product are suitable for use.
- It should be used by experts in the field. (Our products are not used by ordinary people.)
- Check the Scailing factor on the disc.
- Mill the disc into the required framework shapes.
- If coloring liquids are infiltrated into the milled frameworks, dry them for 10-25 min at temperature
lower than 140℃ depending on the size of restorations.
- A close examination of the doctor is required for the patients with bruxism or
hypersensitivity reactions is required.
- Sinter the frameworks in accordance with heating schedule for translucency and strength.
-Design the restoration so that it does not use excessive disc edges.
-The support of the bridge should be supported by pontic as much as possible.
(Not the abutment crown on both sides)
- Qualities can be improved if heat treatment is conducted at up to 1000℃ prior to build-up after sintering has been completed.
- Use the build-up powder for zirconia use as recommended by the manufacturer.
ACSI-20P
Instructions for Use
(Dental Ceramic blocks)
 | Manufacturer |  | Authorized Representative in the European Community
 | Do not reuse |  | Catalogue Number
 | Date of Manufacture |  | Non Sterile
 | Batch Code |  | Keep away from sunlight
 | Consult Instructions for Use |  | Keep dry
 | Use by |  | 
 | Vatechacucera, Inc. C-dong, 2544, Nambuk-daero, Idong-eup, Cheoin-gu, Yongin-si, Gyeonggi-do, Korea |  | Vatech France Floor 4, 51 Quai de Dion Bouton 92800 Puteaux, France
2021.03.03. (Rev.14)
사용설명서1 (안)
Starting temp.(℃) | Target temp. (℃) | Heating rate (℃/min) | Holding time at target temp. (h)
20/ room temp. | 1,200 | 10 | 0
1,200 | Target temp. | 2 | 2
Target temp. | 20/room temp. | Furnace cooling to target temp | 
300x135 양면 [3단접지], 전단 합판 80모조
Sintering
- Recommended Target temp.(℃): Classic POS, Perfit ZR HT (1,550)
Classic UTL, Perfit ZR ST (1,530)/Classic ANT, Perfit ZR UT (1,500)
Post sintering (Precaution about crown fracture)
-Completely sintered restorations can be adjusted through a wet grinding process with a diamond-coated burr using light pressure to prevent overheating and microcracks.
-After sintering, the thickness of the crown should not be thinner than 0.5mm.
In the case of bridges, the cross-sectional area of the connector must be at least 9 ㎟.
-Before veneering, wash and dry the restoration with running water or steam.
-The coefficient of thermal expansion of all zirconia veneering powders should be 10 ~ 10.6x10⁻⁶/℃.
-Use cement for bonding zirconia restorations according to the manufacturer's instructions.
Safety instructions
- Dusts, produced during the preparation of frameworks, may be hazardous to your health.
Processing may therefore be performed only with a properly functioning ventilation system.
Storage method after using - The block after machining is disposable and should not be reused. - Keep the unused block completely dry. - Store the product in its original packaging to prevent damage.
- Storage method: Avoid direct sunlight and store in a dry place at room temperature.
- Do not place where infection is possible
Considerations
- It is recommended to apply for case of single tooth to short bridge(≤3-unit).
* Classic POS and Perfit ZR HT are classified as class 5 can be used as a prosthesis containing
more than 4units.
- Do not load excessively when milling or positioning.
- When processing, please use suitable goggles and protective gloves.
- Milling process techniques require special and complicated procedures.
-Overall understanding should be preceded about application methods, clinical indications and
cautions related to Classic POS / UTL / ANT, Perfit ZR HT / ST / UT.
- If the ceramic powder or liquid come into direct contact with skin and eyes, wash with plenty of
water and consult physician.
- The presence of notches and sharp edges of damage on the bottom side of the inter dental
connectors may substantially reduce the stability of the sintered framework.
- Please do not touch the surface of core with hand after sintering.
- Additional heating of the core sintered in the furnace might cause cracks by deterioration of
mechanical properties.
- Don’t use unsterilized instruments.
- Don’t use without wearing gloves or mask.
- Non-professional use (Use such as dental technicians or dentists)
- Don’t use patients with allergic reactions of the material (zirconia)
- Don’t perform the procedure on patients with inflammation or disease in the mouth
※ Conditions for porcelain powder build up and sintering
- The coefficient of thermal expansion of Zirconia blocks is ~10ppm/℃ and only compatible
Zirconia porcelain powders should be used.
- The thickness of the porcelain powder should be kept under 1mm to reduce its risk of Fracture.
- The sintering temperature of porcelain powders varies according to the manufacturers, please
Follow the manuals for each manufacturer.
- This Medical device is disposable. Prohibit re-use.
Re-use may cause inflammation, diseases (such as infections), breakage and
cracking of the prosthesis.
- It is recommended that the manufactured prosthesis be inspected regularly by a expert.
- This device is required to be disposed as general wastes as per the local regulation etc.
A consideration is required as to dispose the device once it is placed on the teeth.
- Avoid eating hard or tough foods as the product may be damaged broken after
restoration.
Our Products are not subject to the shelf-life date because aging such as changes in
physical properties does not occur during storage.
※ Instruction for Use except for English Version will be translated and provided for each EU member countries after getting an approval on the Instruction for Use from each buyer of the related countries
Notes: 반소결 영문

## Slide 2
ACSI-20P
제품 사용설명서
(절삭가공용치과도재)
2021.04.06. (Rev.13)
사용설명서1 (안)
가. 사용 목적
- 인레이, 인공치, 크라운, 브릿지 등의 치과 수복물을 제작하기 위해서 사용하는 도재로서
치과용 컴퓨터 지원설계, 제조유닛으로 절삭 가공하여 사용한다.
나. 일반적인 취급사항
- 제품의 손상 여부를 확인한다.
- 의도한 용도에 따라 디스크의 종류와 색깔을 확인한다.
- 이물질이나 액체와 접촉하지 않도록 주의한다.
- 블록을 원래의 포장과 건조한 곳에 보관한다.
다. 소결 전
- 제품의 수축률과 확대율을 확인한다.
- 제품을 필요한 모양으로 절삭한다.
- 제조업체에서 권장하는 컬러링 리퀴드를 사용한다.
- 컬러링 리퀴드가 절삭된 수복물에 침투되면 수복물의 크기에 따라 140 ℃ 이하의 온도에서
10-25 분 동안 건조한다.
라. 소결
- 최적의 투명도, 강도를 위해 다음 소결 스케줄에 따른다.
Chart: 
| 0분 | 120분 | 330분 | 550분
온도(℃) | 30 | 1200 | 1500 | 1500 | 100
210분
120분
1200
시간
온도(℃)
- 수복물이 지나치게 블록의 가장자리에서 가공되지 않도록 디자인한다.
- 브릿지의 지지대는 양쪽의 지대치 어버트먼트 크라운에 의해서가 아니라 가급적 폰틱으로
지지되게한다.
- 소결이 완료된 후 빌드업 하기 전에 1,000℃까지 열처리를 하면 품질이 향상된다. - 제조업체에서 권장하는 지르코니아용 빌드업 분말을 사용한다.
최종온도
※ 권장 소결온도(℃): Classic POS, Perfit ZR HT (1,550) / Classic UTL, Perfit ZR ST (1,530)
Classic ANT, Perfit ZR UT (1,500)
마. 소결 후
- 완소결된 수복물은 과열 및 미세 크랙을 방지하기 위해 가벼운 압력을 사용하여 다이아몬드 코팅
버로 습식 분쇄 공정을 통해 조절할 수 있다.
- 소결 후 크라운의 두께가 0.5mm 보다 얇아서는 안된다. 브릿지의 경우 커넥터의 단면적은 9 ㎟
이상이어야 한다.
- 비니어링을 하기 전에, 흐르는 물 또는 스팀으로 수복물을 세척하고 건조한다.
- 모든 지르코니아용 비니어링 분말의 열팽창 계수는 10 ~ 10.6x10⁻⁶/℃ 이어야 한다.
- 지르코니아 수복물 접착용 시멘트는 제조업체의 지침에 따라 사용한다.
바. 사용상 주의사항
- 수복물 가공에 발생하는 분진은 건강상 해로울 수가 있으므로 환기가 원활한 곳이나 분진 포집기
하에서 작업한다.
- 치과의사 및 치기공사 등 전문가를 위한 것으로 일반인의 사용은 불가하다.
- 어린이, 노약자가 다루지 않도록 주의한다.
- 블록의 특정 성분에 알레르기 반응이 있는 환자의 경우는 이 재료를 사용하여 수복물을 제작하지 않는다.
- 바닥에 떨어뜨리거나 충격을 주어 파손되지 않도록 한다.
사. 유효기간
- 보관조건에 따라 저장하였을 경우 반영구적임
아. 기재사항
(1) 제조업자 상호 및 주소
- 상호: ㈜바텍에큐세라
- 주소: 경기도 용인시 처인구 이동읍 남북대로 2544, C동
- 홈페이지: www.acucera.com
(2) 제품명
- 품명: 절삭가공용치과도재
- 모델명: 제품별 별도 기재
(3) 제조번호와 제조년월일: 통관 후 기재
(4) 제조 품목허가번호: 제인10-532호, 제인15-1120호, 제인20-4711호, 제인20-4712호,
제인20-4810호, 제인20-4906호
(5) 사용 목적: 인공 치아, 보철물 등을 만드는 데에 사용되는 재료로 재질은 Zirconium oxide 이다.
(6) 보관 또는 저장방법: 직사광선을 피하고 상온의 건조한 곳에 보관한다.
(7) 본 제품은 비멸균 일회용 의료기기로 재사용 금지
제품
사용설명서
Instruction for Use
(Dental Ceramic Blocks)
표지
맨 뒷면
Notes: 반소결 영문

## Slide 3
27x12 mm 스티커 제작 (10,000장)
마크 가로x세로 비율 유지
1639
↑ 테두리선 인쇄x
스티커 제작